## Slide 1: Optical Wire Position Monitor
Wojciech Żak
wojciech.zak@cern.ch

## Slide 2: Outline
Application of the OPWM in the SPL Mock-up
Idea of the Optical Wire Position Monitor
Conception tests and calibration at warm
Tests at liquid nitrogen
Behavior of components with temperature
Conclusions

## Slide 3
Stretched wire
Opto-coupler sensors
Last cavity support
Courtesy of V. Parma
SPL SCM Cavity position monitoring specs:
Static position or slow movements: absolute movements (x,y,z) of each of 4 cavities during steady state operation and cool-down/warm-ups (300-2 K)
Vertical range 0-2 mm
Precision < 0.05 mm
Resolution < 0.01 mm
Possibly vibration measures (0-1 kHz)

## Slide 4: OWPM position on the SPL SCM
Stretched wire measuring system interface
Thermal shield
interface
Courtesy of CNRS

## Slide 5
How does it work?
wire

## Slide 6
LED and photo-transistor

## Slide 7: Requirements for the sensor
Simple as possible
Work with a known characteristics “at cold” and “at warm”
Adjustable position to cope with a sag from stretched wire
Φ of a wire should have Φ equals to 1/3 of the Φ of a LED
Gap between LED and transistor should be bigger than the Φ of the wire
Minimalize heat load to the cryostat from the light source and receiver
LED and transistor has to be focused (small light angle) and protected from external light sources
Surface of the sensor around the LED and phototransistor should not be reflective
Resistors mounted outside cryostat
Resistor has to be sensitive enough to “capture” voltage drops caused by the temperature drop and caused by wire movement
Equipped with some compensation system to compensate voltage drops caused by cooling down

## Slide 8: Expectations about the first tests
Prove the concept
Characterize LEDs and phototransistors at room temperature

## Slide 9: First concept of the sensor layout
We expect to cover 0.5 mm range with one LED, to fullfill requirement our sensor will be equiped with 4
wire
Electronic circuit for 1 set of components

## Slide 10: Sensor characterization at room temp
The high-precision, linear stages from Newport®, mounted in XY configuration, providing an accuracy of 1 µm on 150 mm maximum displacement. A motor controller, Newport® ESP7000, is used for acquiring the position through a linear encoder. A DC motor with a low-friction ball screw eliminates slip-stick effects and delivers displacements with 100-nm incremental motion capability. Manual adjustments can be made using a knob at the end of the motor.
An important source of the uncertainty is the parallelism of the stages. Using a Leica laser system, based on the optical interferometry effect, and level meters were used to adjust the parallelism.
Courtesy of C. Petrone

## Slide 11: Tests setups
Box to shield the environment light noise,
setup with 2.5mm wire
Test setup with 1.5 mm wire
Courtesy of C. Petrone

## Slide 12
Results
Output voltage signal variations according to the movement of the 2.5mm wire
Courtesy of C. Petrone
Phototransistor 1
Phototransistor 2
Phototransistor 3
Phototransistor 4
2mm!

## Slide 13
Output voltage signal variations according to the movement of the 1.5mm wire
Standard deviation on three measurements in a row
Courtesy of C. Petrone
Phototransistor 1
Phototransistor 2
Phototransistor 3
Phototransistor 4

## Slide 14
Expectations for tests in liquid nitrogen
Try to reproduce plots
Reduce the size of the sensor
Check where to place the resistors: outside or inside the cryostat
Find the best way to power up LEDs and phototransistors
Design PCB which allow us easily change components
Check behavior of the components with temperature

## Slide 15
Setup for tests „at cold”
PCB with two sensors
Courtesy of K. Motala
Cross section of the cryostat
wire
LED/phototransistorr
JUMPER to control which resistors to use

## Slide 16
During the test
Courtesy of K. Motala
Resistors inside the cryostat
Resistors outside the cryostat
77K
77K
300K

## Slide 17
Chart: Sample number
Repetition test, sensor 1 – resistors inside cryostat
Samples were taken every 10 s

## Slide 18
Resistors outside cryostat
Chart: Pumping vacuum and cooling down
Samples were taken every 10 s

## Slide 19
Test at cold with movable wire
New cryostat with movable wire

## Slide 20
Electric circuit:
Adjustable voltage or current supply for the LEDs
Osram LED SFH 484-2
Osram phototransistor SFH313FA
3x PT100: 2 at the bottom, 1 at the top flange
Every sample correspond to 100 microns displacement of the top plate

## Slide 21
First test at cold with movable wire
Chart: Sample number
1 | 4.6423 | 4.5663 | 4.5327 | 4.4567 | 4.3764 | 4.3304 | 4.1676 | 4.048 | 3.9422 | 3.8376 | 3.7124 | 3.5757 | 3.4387 | 3.2945 | 3.1534 | 3.0242 | 2.9027 | 2.7897 | 2.6934 | 2.611 | 2.5497 | 2.5102 | 2.4906 | 2.507 | 2.5269 | 2.5756 | 2.6601 | 2.749 | 2.8443 | 2.9424 | 3.0415 | 3.1478 | 3.2367 | 3.3313 | 3.4299 | 3.533 | 3.6522 | 3.7789 | 3.8818
2 | 4.6499 | 4.5944 | 4.5493 | 4.4887 | 4.3855 | 4.2821 | 4.1792 | 4.0639 | 3.954 | 3.8604 | 3.732 | 3.6194 | 3.4665 | 3.3133 | 3.1708 | 3.0422 | 2.935 | 2.8144 | 2.7217 | 2.632 | 2.5641 | 2.5269 | 2.5018 | 2.5322 | 2.5618 | 2.6123 | 2.6631 | 2.7305 | 2.8297 | 2.9141 | 3.0287 | 3.1374 | 3.2367 | 3.3395 | 3.4367 | 3.5428 | 3.6521 | 3.7622 | 3.8913 | 4.0024 | 4.115
3 | 4.682 | 4.6287 | 4.5665 | 4.5064 | 4.4204 | 4.3486 | 4.2363 | 4.1402 | 4.0201 | 3.8754 | 3.7924 | 3.6345 | 3.4841 | 3.3518 | 3.2092 | 3.0771 | 2.9509 | 2.8448 | 2.7476 | 2.6531 | 2.5917 | 2.5347 | 2.5198 | 2.5319 | 2.5723 | 2.6097 | 2.6586 | 2.7294 | 2.8294 | 2.9182 | 3.0297 | 3.1452 | 3.2424 | 3.3603 | 3.4357 | 3.5561 | 3.6295 | 3.7549 | 3.8986 | 3.996 | 4.1655
c1 | 4.0003 | 3.9355 | 3.8672 | 3.8008 | 3.7405 | 3.657 | 3.5637 | 3.4704 | 3.3631 | 3.2188 | 3.0879 | 2.8881 | 2.6964 | 2.5183 | 2.3503 | 2.2128 | 2.0747 | 1.9642 | 1.9015 | 1.8463 | 1.825 | 1.8093 | 1.8169 | 1.8481 | 1.8868 | 1.9337 | 2.0151 | 2.1204 | 2.2489 | 2.3841 | 2.5488 | 2.7035 | 2.8583 | 2.992 | 3.1041 | 3.1979 | 3.2872
c2 | 4.8881 | 4.877 | 4.8628 | 4.8651 | 4.8077 | 4.7478 | 4.7047 | 4.5642 | 4.4201 | 4.3146 | 4.1563 | 4.0228 | 3.8762 | 3.7293 | 3.6303 | 3.487 | 3.3881 | 3.2462 | 3.1826 | 3.0794 | 2.9594 | 2.855 | 2.797 | 2.775 | 2.7596 | 2.776 | 2.8122 | 2.7772 | 2.8339 | 2.9171 | 3.0195 | 3.1363 | 3.262 | 3.4042 | 3.5068 | 3.6529 | 3.796 | 3.9379 | 4.0591 | 4.2179 | 4.3224 | 4.4889 | 4.5919 | 4.6733 | 4.7674

## Slide 22
Lists of improvements:
New length of the cryostat (0.5m) to have bigger temperature drop
Fisher plugs
PCB to easily change LED and/or phototransistors
Fixation for the wire
Upgraded thermalisation for the cables
Added:
Clamps for test cryostat to the „tank” for liquid N2
Support to better guide bar with the wire
2 more linear potentiometers

## Slide 23
The new PCB to test different LEDs/phototransistors

## Slide 24
New guidance for the bar with the wire and potentiometers

## Slide 25
Behavior of phototransistor with LED powered up with constant current
Chart: sample number
1c | 4.1921 | 4.1727 | 4.2277 | 4.1286 | 4.1023 | 4.0696 | 4.0241 | 3.9821 | 3.9247 | 3.8728 | 3.818 | 3.7042 | 3.6094 | 3.5121 | 3.4413 | 3.337 | 3.1868 | 3.0527 | 2.9227 | 2.7779 | 2.6167 | 2.4634 | 2.3179 | 2.1844 | 2.0959 | 2.0013 | 1.9485 | 1.9121 | 1.8958 | 1.9038 | 1.9216 | 1.9572 | 2.0092 | 2.094 | 2.1662 | 2.2341 | 2.3719 | 2.5224 | 2.6562 | 2.8131 | 2.943 | 3.0761 | 3.1931 | 3.2984 | 3.3916 | 3.4789 | 3.5352 | 3.6203 | 3.7175 | 3.7889 | 3.8752 | 3.9379 | 3.9929 | 4.0289 | 4.0414 | 4.0633 | 4.0861
2c | 4.1429 | 4.1271 | 4.1193 | 4.0679 | 4.0554 | 4.0102 | 3.9694 | 3.9236 | 3.8721 | 3.8175 | 3.7375 | 3.6553 | 3.5638 | 3.4803 | 3.3822 | 3.263 | 3.1495 | 3.0447 | 2.8827 | 2.7404 | 2.586 | 2.4348 | 2.2829 | 2.1664 | 2.0573 | 1.9805 | 1.9248 | 1.8836 | 1.8768 | 1.884 | 1.889 | 1.9188 | 1.9764 | 2.0383 | 2.1231 | 2.2065 | 2.3255 | 2.459 | 2.6112 | 2.7552 | 2.888 | 3.0239 | 3.1284 | 3.211 | 3.2956 | 3.3882 | 3.4711 | 3.4776 | 3.6255 | 3.6951 | 3.7902 | 3.857 | 3.8828 | 3.9173 | 3.9357 | 3.9491 | 3.9782
3c | 4.1921 | 4.1727 | 4.2277 | 4.1286 | 4.1023 | 4.0696 | 4.0241 | 3.9821 | 3.9247 | 3.8728 | 3.818 | 3.7042 | 3.6094 | 3.5121 | 3.4413 | 3.337 | 3.1868 | 3.0527 | 2.9227 | 2.7779 | 2.6167 | 2.4634 | 2.3179 | 2.1844 | 2.0959 | 2.0013 | 1.9485 | 1.9121 | 1.8958 | 1.9038 | 1.9216 | 1.9572 | 2.0092 | 2.094 | 2.1662 | 2.2341 | 2.3719 | 2.5224 | 2.6562 | 2.8131 | 2.943 | 3.0761 | 3.1931 | 3.2984 | 3.3916 | 3.4789 | 3.5352 | 3.6203 | 3.7175 | 3.7889 | 3.8752 | 3.9379 | 3.9929 | 4.0289 | 4.0414 | 4.0633 | 4.0861
1w | 3.9755 | 3.9885 | 3.9573 | 3.9187 | 3.8967 | 3.8685 | 3.8377 | 3.8041 | 3.7568 | 3.7002 | 3.6568 | 3.607 | 3.5117 | 3.4318 | 3.3492 | 3.2478 | 3.1405 | 3.0031 | 2.8587 | 2.7363 | 2.5822 | 2.4369 | 2.281 | 2.1677 | 2.0582 | 1.9856 | 1.9221 | 1.8909 | 1.875 | 1.8766 | 1.9063 | 1.9212 | 1.9621 | 2.0286 | 2.1191 | 2.2565 | 2.3797 | 2.541 | 2.6602 | 2.8007 | 2.9465 | 3.0565 | 3.1514 | 3.2366 | 3.3177 | 3.4037 | 3.5006 | 3.5828 | 3.6699 | 3.7622 | 3.8253 | 3.8938 | 3.9557 | 3.9826 | 4.0012 | 4.0202
2w | 3.9755 | 3.9885 | 3.9573 | 3.9187 | 3.8967 | 3.8685 | 3.8377 | 3.8041 | 3.7568 | 3.7002 | 3.6568 | 3.607 | 3.5117 | 3.4318 | 3.3492 | 3.2478 | 3.1405 | 3.0031 | 2.8587 | 2.7363 | 2.5822 | 2.4369 | 2.281 | 2.1677 | 2.0582 | 1.9856 | 1.9221 | 1.8909 | 1.875 | 1.8766 | 1.9063 | 1.9212 | 1.9621 | 2.0286 | 2.1191 | 2.2565 | 2.3797 | 2.541 | 2.6602 | 2.8007 | 2.9465 | 3.0565 | 3.1514 | 3.2366 | 3.3177 | 3.4037 | 3.5006 | 3.5828 | 3.6699 | 3.7622 | 3.8253 | 3.8938 | 3.9557 | 3.9826 | 4.0012 | 4.0202
3w | 3.9705 | 3.9579 | 3.9239 | 3.896 | 3.8916 | 3.8423 | 3.8139 | 3.7744 | 3.7185 | 3.6605 | 3.6177 | 3.527 | 3.4511 | 3.3769 | 3.2654 | 3.1486 | 3.0061 | 2.9007 | 2.7746 | 2.7152 | 2.6147 | 2.4691 | 2.3119 | 2.2075 | 2.0922 | 1.9934 | 1.9367 | 1.8983 | 1.8794 | 1.877 | 1.8824 | 1.9165 | 1.951 | 2.0126 | 2.1029 | 2.1673 | 2.3055 | 2.4182 | 2.5788 | 2.7458 | 2.8809 | 2.994 | 3.1054 | 3.1945 | 3.2784 | 3.3662 | 3.4811 | 3.5618 | 3.6444 | 3.7383 | 3.8311 | 3.8691 | 3.9247 | 3.9706 | 3.9911 | 4.0216
w stands for „warm” (300K)
c stands for „cold” (77K)

## Slide 26
Chart: sample number
1c | 4.843 | 4.8209 | 4.8241 | 4.8146 | 4.7946 | 4.7891 | 4.7424 | 4.7134 | 4.6444 | 4.5734 | 4.4925 | 4.4005 | 4.3464 | 4.2289 | 4.1146 | 3.9974 | 3.8454 | 3.6897 | 3.5051 | 3.2832 | 3.1008 | 2.918 | 2.7626 | 2.6216 | 2.5089 | 2.427 | 2.3748 | 2.3468 | 2.3286 | 2.3219 | 2.3358 | 2.3622 | 2.4228 | 2.4832 | 2.5749 | 2.6732 | 2.8092 | 2.9349 | 3.0847 | 3.2242 | 3.3824 | 3.5158 | 3.6429 | 3.7428 | 3.8285 | 3.896 | 3.9791 | 4.0541 | 4.1295 | 4.2177 | 4.3072 | 4.4024 | 4.5008 | 4.5869 | 4.6557 | 4.6841 | 4.7149
2c | 4.8571 | 4.8579 | 4.8513 | 4.8496 | 4.8323 | 4.8248 | 4.8004 | 4.7592 | 4.7012 | 4.6267 | 4.5679 | 4.4911 | 4.3947 | 4.2904 | 4.1887 | 4.0613 | 3.9199 | 3.7664 | 3.5995 | 3.3686 | 3.1481 | 2.9689 | 2.7899 | 2.6599 | 2.5592 | 2.4905 | 2.4383 | 2.3927 | 2.3803 | 2.3785 | 2.389 | 2.4216 | 2.4823 | 2.5519 | 2.6747 | 2.7581 | 2.8886 | 3.0329 | 3.1792 | 3.3449 | 3.4826 | 3.6169 | 3.7551 | 3.8527 | 3.9383 | 4.0359 | 4.1038 | 4.1931 | 4.2867 | 4.3499 | 4.4862 | 4.5294 | 4.6453 | 4.7049 | 4.7491 | 4.7751 | 4.7987
3c | 4.8899 | 4.8829 | 4.8835 | 4.879 | 4.8751 | 4.8629 | 4.8514 | 4.8318 | 4.7759 | 4.7306 | 4.6735 | 4.5909 | 4.5004 | 4.4019 | 4.3278 | 4.1942 | 4.0364 | 3.8599 | 3.6593 | 3.4424 | 3.2526 | 3.0409 | 2.8649 | 2.7276 | 2.6293 | 2.527 | 2.5009 | 2.4699 | 2.4441 | 2.4533 | 2.4568 | 2.5063 | 2.57 | 2.6562 | 2.7534 | 2.8305 | 2.987 | 3.1287 | 3.2678 | 3.4423 | 3.5909 | 3.7571 | 3.8754 | 3.9866 | 4.0878 | 4.1644 | 4.2409 | 4.3276 | 4.4233 | 4.5154 | 4.6046 | 4.6928 | 4.7743 | 4.7967 | 4.8197 | 4.8393 | 4.8512
1w | 4.8232 | 4.8027 | 4.8489 | 4.8099 | 4.8048 | 4.7774 | 4.7286 | 4.6942 | 4.6434 | 4.5679 | 4.5051 | 4.4193 | 4.3527 | 4.282 | 4.1923 | 4.0745 | 3.9216 | 3.7758 | 3.5947 | 3.4183 | 3.2017 | 3.0268 | 2.8806 | 2.7719 | 2.6666 | 2.6177 | 2.5659 | 2.5344 | 2.5256 | 2.5298 | 2.5672 | 2.5895 | 2.6572 | 2.7305 | 2.8327 | 2.9578 | 3.1097 | 3.2622 | 3.3574 | 3.5105 | 3.6473 | 3.7555 | 3.8626 | 3.9565 | 4.0362 | 4.0978 | 4.1879 | 4.2593 | 4.3354 | 4.4285 | 4.5177 | 4.5975 | 4.6722 | 4.738 | 4.7822 | 4.8022
2w | 4.8175 | 4.776 | 4.7905 | 4.7748 | 4.7297 | 4.7107 | 4.6484 | 4.5988 | 4.5333 | 4.4654 | 4.4109 | 4.3482 | 4.2855 | 4.2029 | 4.1154 | 4.0082 | 3.8596 | 3.7044 | 3.5114 | 3.3031 | 3.1434 | 2.9646 | 2.817 | 2.7133 | 2.6345 | 2.572 | 2.5273 | 2.4994 | 2.49 | 2.4958 | 2.5262 | 2.5578 | 2.6121 | 2.7011 | 2.8175 | 2.9345 | 3.0464 | 3.1625 | 3.3045 | 3.439 | 3.5813 | 3.6894 | 3.7565 | 3.8388 | 3.9378 | 4.0034 | 4.0827 | 4.1821 | 4.2411 | 4.3131 | 4.4066 | 4.4883 | 4.5683 | 4.6533 | 4.753 | 4.7619
3w | 4.7783 | 4.7811 | 4.7776 | 4.7548 | 4.716 | 4.6751 | 4.6332 | 4.5694 | 4.5024 | 4.4418 | 4.369 | 4.304 | 4.2405 | 4.1513 | 4.0368 | 3.9012 | 3.7135 | 3.5647 | 3.3901 | 3.3117 | 3.1783 | 3.0392 | 2.8358 | 2.7458 | 2.649 | 2.5711 | 2.5211 | 2.4905 | 2.4915 | 2.4864 | 2.5018 | 2.5361 | 2.5811 | 2.6622 | 2.7667 | 2.8398 | 2.9708 | 3.0888 | 3.2292 | 3.3613 | 3.5151 | 3.6362 | 3.7324 | 3.8287 | 3.9423 | 4.016 | 4.0959 | 4.18 | 4.2433 | 4.3487 | 4.4236 | 4.475 | 4.5648 | 4.6556 | 4.7339 | 4.7646
Behavior of phototransistor with LED powered up with constant voltage drop
w stands for „warm” (300K)
c stands for „cold” (77K)

## Slide 27
LED
RECEIVER
PT100
300K
PT100
77K
Behavior of components with temperature

## Slide 28
LED
RECEIVER
300K
77K
Temperature during tests
Voltage drop at the LED
Voltage drop at the phototransistor 2
Chart: Temperature [K]
Chart: Time [sample]
Chart: Temperature [K]

## Slide 29: Zoom at the “77” K, phototransistor
Chart: Temperature [K]
Chart: Temperature [K]
Phototransistor 1 in time

## Slide 30
Chart: Time
Background measurements

## Slide 31: Phototransistor 1: background measurement with the LED on (t=3h)
Chart: sample number
Test condition:
290-294K
More than 12h after test in liquid N2

## Slide 32: Phototransistor 2: background measurement with the LED on (t=3h)
Chart: sample number
Test condition:
290-294K
More than 12h after test in liquid N2

## Slide 33
The biggest problem is lack of linear reaction to the temperature and complete different behavior of components with temperature: LED is giving more light when phototransistor is becoming less sensitive. This makes it almost impossible to apply for the transient conditions.
During the tests we’ve notice different behavior of components strongly dependent on temperature: the same test condition and different results (starting and ending point).
Final sensor will have to be equipped with a temperature sensor (for example PT100).
We can easily assume that one measurement point (in X and Y axes) will require around 10 cables: power/ground for phototransistors (3-4 cables), power/ground to the LEDs (3-4), temperature sensor (4 cables) and possibly 3 more cables for feedback system (led + phototransistor). Just for the mock-up we have 8 targets for the laser tracker: if we replace them with our solution it will make almost 80 cables for just one cavity.
Conclusions:

## Slide 34
Thank you for your attention!
&
Big „thank you”: to A. Vande Craen, K. Motala, M. Zapolski, M. Krupa, M. Strychalski, R. Bonnomi,
M. Souchet, X. Sauvage, A. Guimet, A. Bastard, C. Petrone, J. G. Perez,